## Sheet1
DRIFT ANALYSIS OF MULTIFUNCTION CALIBRATOR(5520A) | [17 empty] | Within/out of permissible error/acceptable limit
| UUC value |  | Std. value | UUC value |  | Std value | Error (cal point - reading 2019) | Error (cal point - reading 2021) | Error (cal point - reading 2021) |  | Trend of result (line graph) | absolute drift from mean | abolute drift from mean | Trend of result (bar graph) | 2019-2021 (based on absolute drift from mean) | Eroded parameters | Permissible error | Status (2019) | Status (2021)
Parameter | 2019 |  | 2019 | 2021 |  | 2021 | 2019 | 2021 | 2021 |  | Error 2019 to error 2021 | 2019 | 2021 | Mean drift 2019-2021 | reduced drift = improvement same drift = same cond. increased drift = degradation
DCV | 0.1 | mV | 0.09912 | 0.1 | mV | 0.09969 | 0.00088 | 0.00031 | 3.1e-07 | V |  | 0.00088 | 0.00031 |  | improved |  | 1.2e-06 | out of limit | out of limit
DCV | 1 | mV | 0.99917 | 1 | mV | 0.99962 | 0.00083 | 0.00038 | 3.8e-07 | V |  | 0.00083 | 0.00038 |  | improved |  | 1.2e-05 | out of limit | out of limit | improved | 186 | Degraded parameters | Nos. of degradation
DCV | 10 | mV | 9.9991 | 10 | mV | 0.99965 | 0.0009 | 9.0003 | 0.0090004 | V |  | 0.0009 | 9.0003 |  | degradation | DCV | 0.00012 | out of limit | out of limit | degradation | 149
DCV | 100 | mV | 99.9984 | 100 | mV | 99.9998 | 0.00159 | 0.00024 | 2.4e-07 | V |  | 0.00159 | 0.00024 |  | improved |  | 0.0012 | out of limit | within limit |  | 0 | #REF!
DCV | 329 | mV | 328.9979 | 329 | mV | 328.9999 | 0.00212 | 0.0001 | 1e-07 | V |  | 0.00212 | 0.0001 |  | improved |  | 0.003948 | within limit | within limit |  |  |  | 0
DCV | 1 | V | 0.99999 | 1 | V | 1 | 5.7e-06 | -3.9e-05 | -3.9e-05 | V |  | 5.7e-06 | 3.9e-05 |  | improved |  | 1.2e-05 | within limit | out of limit | 2021 |  |  | 0
DCV | 3.29 | V | 3.29 | 3.29 | V | 3.29 | 1.83e-05 | 2e-06 | 2e-06 | V |  | 1.83e-05 | 2e-06 |  | improved |  | 3.948e-05 | within limit | within limit | out of limit | 111 |  | 0
DCV | 10 | V | 10 | 10 | V | 10 | 5e-05 | -1.7e-05 | -1.7e-05 | V |  | 5e-05 | 1.7e-05 |  | improved |  | 0.00012 | within limit | within limit | within limit | 202 |  | 0
DCV | 32.9 | V | 32.8998 | 32.9 | V | 32.8999 | 0.000174 | 0.0001 | 0.0001 | V |  | 0.000174 | 0.0001 |  | improved |  | 0.0003948 | within limit | within limit |  |  |  | 0
DCV | 50 | V | 49.9995 | 50 | V | 50.0001 | 0.00045 | -0.00012 | -0.00012 | V |  | 0.00045 | 0.00012 |  | improved |  | 0.0006 | within limit | within limit | 2019 |  |  | 0
DCV | 100 | V | 99.9995 | 100 | V | 99.9998 | 0.00053 | 0.00015 | 0.00015 | V |  | 0.00053 | 0.00015 |  | improved |  | 0.0012 | within limit | within limit | out of limit | 165 |  | 0
DCV | 329 | V | 328.9981 | 329 | V | 328.9983 | 0.0019 | 0.0017 | 0.0017 | V |  | 0.0019 | 0.0017 |  | improved |  | 0.003948 | within limit | within limit | within limit | 182 |  | 0
DCV | 500 | V | 499.9968 | 500 | V | 499.9984 | 0.0032 | 0.0016 | 0.0016 | V |  | 0.0032 | 0.0016 |  | improved |  | 0.006 | within limit | within limit |  |  | Total | 0
DCV | 1000 | V | 999.9964 | 1000 | V | 999.9974 | 0.0036 | 0.0026 | 0.0026 | V |  | 0.0036 | 0.0026 |  | improved |  | 0.012 | within limit | within limit
ACV @ 50Hz | 1 | mV | 1.004 | 1 | mV | 0.9968 | -0.004 | 0.0032 | 3.2e-06 | V |  | 0.004 | 0.0032 |  | degradation | ACV @ 50Hz | 0.00012 | out of limit | out of limit
ACV @ 50Hz | 10 | mV | 10.0094 | 10 | mV | 9.9962 | -0.0094 | 0.0038 | 3.8e-06 | V |  | 0.0094 | 0.0038 |  | degradation | ACV @ 50Hz | 0.0012 | out of limit | out of limit
ACV @ 50Hz | 32 | mV | 32.0041 | 32 | mV | 31.9953 | -0.0041 | 0.0047 | 4.7e-06 | V |  | 0.0041 | 0.0047 |  | degradation | ACV @ 50Hz | 0.00384 | out of limit | out of limit
ACV @ 50Hz | 50 | mV | 50.0043 | 50 | mV | 49.9942 | -0.0043 | 0.0058 | 5.8e-06 | V |  | 0.0043 | 0.0058 |  | degradation | ACV @ 50Hz | 0.006 | within limit | within limit
ACV @ 50Hz | 100 | mV | 100.0029 | 100 | mV | 99.9981 | -0.0029 | 0.0019 | 1.9e-06 | V |  | 0.0029 | 0.0019 |  | degradation | ACV @ 50Hz | 0.012 | within limit | within limit
ACV @ 50Hz | 329 | mV | 328.997 | 329 | mV | 328.997 | 0.003 | 0.003 | 3e-06 | V |  | 0.003 | 0.003 |  | improved |  | 0.03948 | within limit | within limit
ACV @ 50Hz | 1 | V | 1 | 1 | V | 1.0001 | -1e-06 | -5.8e-05 | -5.8e-05 | V |  | 1e-06 | 5.8e-05 |  | improved |  | 0.00012 | within limit | within limit
ACV @ 50Hz | 3.2 | V | 3.2 | 3.2 | V | 3.2003 | 2e-05 | -0.00034 | -0.00034 | V |  | 2e-05 | 0.00034 |  | improved |  | 0.000384 | within limit | within limit
ACV @ 50Hz | 10 | V | 10.0001 | 10 | V | 10.0003 | -0.00013 | -0.00028 | -0.00028 | V |  | 0.00013 | 0.00028 |  | improved |  | 0.0012 | within limit | within limit
ACV @ 50Hz | 32 | V | 32.0001 | 32 | V | 31.9999 | -0.0001 | 0.0001 | 0.0001 | V |  | 0.0001 | 0.0001 |  | degradation | ACV @ 50Hz | 0.00384 | within limit | within limit
ACV @ 50Hz | 50 | V | 50.0015 | 50 | V | 50.0077 | -0.0015 | -0.0077 | -0.0077 | V |  | 0.0015 | 0.0077 |  | improved |  | 0.006 | within limit | out of limit
ACV @ 50Hz | 100 | V | 100.0018 | 100 | V | 100.0065 | -0.0018 | -0.0065 | -0.0065 | V |  | 0.0018 | 0.0065 |  | improved |  | 0.012 | within limit | within limit
ACV @ 50Hz | 329 | V | 329.015 | 329 | V | 329.008 | -0.015 | -0.008 | -0.008 | V |  | 0.015 | 0.008 |  | degradation | ACV @ 50Hz | 0.03948 | within limit | within limit
ACV @ 50Hz | 500 | V | 500.014 | 500 | V | 500.023 | -0.014 | -0.023 | -0.023 | V |  | 0.014 | 0.023 |  | improved |  | 0.06 | within limit | within limit
ACV @ 50Hz | 1000 | V | 1000.078 | 1000 | V | 1000.048 | -0.078 | -0.048 | -0.048 | V |  | 0.078 | 0.048 |  | degradation | ACV @ 50Hz | 0.12 | within limit | within limit
ACV @ 1kHz | 1 | mV | 0.9993 | 1 | mV | 0.996 | 0.0007 | 0.004 | 4e-06 | V |  | 0.0007 | 0.004 |  | improved |  | 0.00012 | out of limit | out of limit
ACV @ 1kHz | 10 | mV | 10.0094 | 10 | mV | 9.995 | -0.0094 | 0.005 | 5e-06 | V |  | 0.0094 | 0.005 |  | degradation | ACV @ 1kHz | 0.0012 | out of limit | out of limit
ACV @ 1kHz | 32 | mV | 32.0043 | 32 | mV | 31.9974 | -0.0043 | 0.0026 | 2.6e-06 | V |  | 0.0043 | 0.0026 |  | degradation | ACV @ 1kHz | 0.00384 | out of limit | within limit
ACV @ 1kHz | 50 | mV | 50.0052 | 50 | mV | 49.9961 | -0.0052 | 0.0039 | 3.9e-06 | V |  | 0.0052 | 0.0039 |  | degradation | ACV @ 1kHz | 0.006 | within limit | within limit
ACV @ 1kHz | 100 | mV | 100.0031 | 100 | mV | 100.008 | -0.0031 | -0.008 | -8e-06 | V |  | 0.0031 | 0.008 |  | degradation | ACV @ 1kHz | 0.012 | within limit | within limit
ACV @ 1kHz | 329 | mV | 328.999 | 329 | mV | 329.007 | 0.001 | -0.007 | -7e-06 | V |  | 0.001 | 0.007 |  | improved |  | 0.03948 | within limit | within limit
ACV @ 1kHz | 1 | V | 1 | 1 | V | 1.0001 | -6e-06 | -0.000131 | -0.000131 | V |  | 6e-06 | 0.000131 |  | improved |  | 0.00012 | within limit | out of limit
ACV @ 1kHz | 3.2 | V | 3.2 | 3.2 | V | 3.2003 | -4e-06 | -0.00032 | -0.00032 | V |  | 4e-06 | 0.00032 |  | improved |  | 0.000384 | within limit | within limit
ACV @ 1kHz | 10 | V | 9.9997 | 10 | V | 10.0011 | 0.00031 | -0.00108 | -0.00108 | V |  | 0.00031 | 0.00108 |  | improved |  | 0.0012 | within limit | within limit
ACV @ 1kHz | 32 | V | 31.9986 | 32 | V | 32.0005 | 0.0014 | -0.0005 | -0.0005 | V |  | 0.0014 | 0.0005 |  | improved |  | 0.00384 | within limit | within limit
ACV @ 1kHz | 50 | V | 49.9991 | 50 | V | 50.0097 | 0.0009 | -0.0097 | -0.0097 | V |  | 0.0009 | 0.0097 |  | improved |  | 0.006 | within limit | out of limit
ACV @ 1kHz | 100 | V | 99.997 | 100 | V | 100.0091 | 0.003 | -0.0091 | -0.0091 | V |  | 0.003 | 0.0091 |  | improved |  | 0.012 | within limit | within limit
ACV @ 1kHz | 329 | V | 328.987 | 329 | V | 329.025 | 0.013 | -0.025 | -0.025 | V |  | 0.013 | 0.025 |  | improved |  | 0.03948 | within limit | within limit
ACV @ 1kHz | 500 | V | 500.012 | 500 | V | 500.078 | -0.012 | -0.078 | -0.078 | V |  | 0.012 | 0.078 |  | improved |  | 0.06 | within limit | out of limit
ACV @ 1kHz | 1000 | V | 1000.03 | 1000 | V | 1000.107 | -0.03 | -0.107 | -0.107 | V |  | 0.03 | 0.107 |  | improved |  | 0.12 | within limit | within limit
ACV @ 20kHz | 1 | mV | 0.9987 | 1 | mV | 0.9951 | 0.0013 | 0.0049 | 4.9e-06 | V |  | 0.0013 | 0.0049 |  | improved |  | 0.00012 | out of limit | out of limit
ACV @ 20kHz | 10 | mV | 10.0094 | 10 | mV | 9.9951 | -0.0094 | 0.0049 | 4.9e-06 | V |  | 0.0094 | 0.0049 |  | degradation | ACV @ 20kHz | 0.0012 | out of limit | out of limit
ACV @ 20kHz | 32 | mV | 32.0062 | 32 | mV | 31.9947 | -0.0062 | 0.0053 | 5.3e-06 | V |  | 0.0062 | 0.0053 |  | degradation | ACV @ 20kHz | 0.00384 | out of limit | out of limit
ACV @ 20kHz | 50 | mV | 50.005 | 50 | mV | 49.9936 | -0.005 | 0.0064 | 6.4e-06 | V |  | 0.005 | 0.0064 |  | degradation | ACV @ 20kHz | 0.006 | within limit | out of limit
ACV @ 20kHz | 100 | mV | 100.0026 | 100 | mV | 99.999 | -0.0026 | 0.001 | 1e-06 | V |  | 0.0026 | 0.001 |  | degradation | ACV @ 20kHz | 0.012 | within limit | within limit
ACV @ 20kHz | 329 | mV | 328.992 | 329 | mV | 329.001 | 0.008 | -0.001 | -1e-06 | V |  | 0.008 | 0.001 |  | improved |  | 0.03948 | within limit | within limit
ACV @ 20kHz | 1 | V | 1 | 1 | V | 1.0001 | -4.2e-05 | -0.000139 | -0.000139 | V |  | 4.2e-05 | 0.000139 |  | improved |  | 0.00012 | within limit | out of limit
ACV @ 20kHz | 3.2 | V | 3.2001 | 3.2 | V | 3.2004 | -6e-05 | -0.00038 | -0.00038 | V |  | 6e-05 | 0.00038 |  | improved |  | 0.000384 | within limit | within limit
ACV @ 20kHz | 10 | V | 9.9996 | 10 | V | 10.0008 | 0.00037 | -0.00075 | -0.00075 | V |  | 0.00037 | 0.00075 |  | improved |  | 0.0012 | within limit | within limit
ACV @ 20kHz | 32 | V | 31.9998 | 32 | V | 32.0008 | 0.0002 | -0.0008 | -0.0008 | V |  | 0.0002 | 0.0008 |  | improved |  | 0.00384 | within limit | within limit
ACV @ 20kHz | 50 | V | 49.993 | 50 | V | 50.0033 | 0.007 | -0.0033 | -0.0033 | V |  | 0.007 | 0.0033 |  | improved |  | 0.006 | out of limit | within limit
ACV @ 20kHz | 100 | V | 99.9855 | 100 | V | 99.9942 | 0.0145 | 0.0058 | 0.0058 | V |  | 0.0145 | 0.0058 |  | improved |  | 0.012 | out of limit | within limit
DCA | 10 | µA | 10.0005 | 10 | µA | 10.0004 | -0.0005 | -0.00036 | -3.6e-10 | A |  | 0.0005 | 0.00036 |  | degradation | DCA | 0.001 | within limit | within limit
DCA | 100 | µA | 99.9988 | 100 | µA | 100.0007 | 0.0012 | -0.00069 | -6.9e-10 | A |  | 0.0012 | 0.00069 |  | improved |  | 0.01 | within limit | within limit
DCA | 190 | µA | 189.9972 | 190 | µA | 190.0007 | 0.0028 | -0.00071 | -7.1e-10 | A |  | 0.0028 | 0.00071 |  | improved |  | 0.019 | within limit | within limit
DCA | 329 | µA | 328.9939 | 329 | µA | 329.0104 | 0.0061 | -0.0104 | -1.04e-08 | A |  | 0.0061 | 0.0104 |  | improved |  | 0.0329 | within limit | within limit
DCA | 0.5 | mA | 0.49999 | 0.5 | mA | 0.5 | 1.2e-05 | 2.5e-06 | 2.5e-09 | A |  | 1.2e-05 | 2.5e-06 |  | improved |  | 5e-05 | within limit | within limit
DCA | 1 | mA | 0.99999 | 1 | mA | 1 | 6e-06 | -3e-07 | -3e-10 | A |  | 6e-06 | 3e-07 |  | improved |  | 0.0001 | within limit | within limit
DCA | 3.29 | mA | 3.29 | 3.29 | mA | 3.29 | 2.6e-05 | -2.6e-05 | -2.6e-08 | A |  | 2.6e-05 | 2.6e-05 |  | improved |  | 0.000329 | within limit | within limit
DCA | 10 | mA | 10.0001 | 10 | mA | 10.0001 | -0.00012 | -0.000137 | -1.37e-07 | A |  | 0.00012 | 0.000137 |  | degradation |  | 0.001 | within limit | within limit
DCA | 20 | mA | 20.0001 | 20 | mA | 19.9999 | -0.00014 | 0.0001 | 1e-07 | A |  | 0.00014 | 0.0001 |  | degradation |  | 0.002 | within limit | within limit
DCA | 32.9 | mA | 32.9002 | 32.9 | mA | 32.8999 | -0.00016 | 0.00014 | 1.4e-07 | A |  | 0.00016 | 0.00014 |  | degradation | DCA | 0.00329 | within limit | within limit
DCA | 100 | mA | 100.0051 | 100 | mA | 100.0006 | -0.00508 | -0.00065 | -6.5e-07 | A |  | 0.00508 | 0.00065 |  | degradation | DCA | 0.01 | within limit | within limit
DCA | 200 | mA | 200.0102 | 200 | mA | 199.9985 | -0.01021 | 0.0015 | 1.5e-06 | A |  | 0.01021 | 0.0015 |  | degradation | DCA | 0.02 | within limit | within limit
DCA | 329.9 | mA | 329.9159 | 329.9 | mA | 329.8915 | -0.01592 | 0.0085 | 8.5e-06 | A |  | 0.01592 | 0.0085 |  | degradation | DCA | 0.03299 | within limit | within limit
DCA | 0.5 | A | 0.50009 | 0.5 | A | 0.49999 | -8.6e-05 | 8.8e-06 | 8.8e-06 | A |  | 8.6e-05 | 8.8e-06 |  | degradation | DCA | 5e-05 | out of limit | within limit
DCA | 1 | A | 1.0002 | 1 | A | 1 | -0.000209 | -2.5e-05 | -2.5e-05 | A |  | 0.000209 | 2.5e-05 |  | degradation | DCA | 0.0001 | out of limit | within limit
DCA | 2.999 | A | 2.9992 | 2.999 | A | 3 | -0.000208 | -0.000972 | -0.000972 | A |  | 0.000208 | 0.000972 |  | improved |  | 0.0002999 | within limit | out of limit
DCA | 5 | A | 4.9991 | 5 | A | 4.9995 | 0.000897 | 0.000465 | 0.000465 | A |  | 0.000897 | 0.000465 |  | improved |  | 0.0005 | out of limit | within limit
DCA | 10 | A | 9.9979 | 10 | A | 9.9978 | 0.002072 | 0.00216 | 0.00216 | A |  | 0.002072 | 0.00216 |  | degradation | DCA | 0.001 | out of limit | out of limit
DCA | 20 | A | 19.9967 | 20 | A | 20.0014 | 0.003281 | -0.00142 | -0.00142 | A |  | 0.003281 | 0.00142 |  | improved |  | 0.002 | out of limit | within limit
ACA @ 50Hz | 33 | µA | 33.011 | 33 | µA | 32.996 | -0.011 | 0.004 | 4e-09 | A |  | 0.011 | 0.004 |  | degradation | ACA @ 50Hz | 0.0198 | within limit | within limit
ACA @ 50Hz | 100 | µA | 100.041 | 100 | µA | 100.0075 | -0.041 | -0.0075 | -7.5e-09 | A |  | 0.041 | 0.0075 |  | degradation | ACA @ 50Hz | 0.06 | within limit | within limit
ACA @ 50Hz | 200 | µA | 200.035 | 200 | µA | 200.015 | -0.035 | -0.015 | -1.5e-08 | A |  | 0.035 | 0.015 |  | degradation | ACA @ 50Hz | 0.12 | within limit | within limit
ACA @ 50Hz | 329 | µA | 329.069 | 329 | µA | 329.04 | -0.069 | -0.04 | -4e-08 | A |  | 0.069 | 0.04 |  | degradation | ACA @ 50Hz | 0.1974 | within limit | within limit
ACA @ 50Hz | 1 | mA | 0.99996 | 1 | mA | 1.0001 | 3.6e-05 | -8.4e-05 | -8.4e-08 | A |  | 3.6e-05 | 8.4e-05 |  | improved |  | 0.0006 | within limit | within limit
ACA @ 50Hz | 3.29 | mA | 3.2909 | 3.29 | mA | 3.29 | -0.000878 | 3e-05 | 3e-08 | A |  | 0.000878 | 3e-05 |  | degradation | ACA @ 50Hz | 0.001974 | within limit | within limit
ACA @ 50Hz | 10 | mA | 10.0029 | 10 | mA | 10.0017 | -0.00294 | -0.00169 | -1.69e-06 | A |  | 0.00294 | 0.00169 |  | degradation | ACA @ 50Hz | 0.006 | within limit | within limit
ACA @ 50Hz | 32.9 | mA | 32.9093 | 32.9 | mA | 32.9021 | -0.00926 | -0.0021 | -2.1e-06 | A |  | 0.00926 | 0.0021 |  | degradation | ACA @ 50Hz | 0.01974 | within limit | within limit
ACA @ 50Hz | 100 | mA | 100.031 | 100 | mA | 100.0177 | -0.031 | -0.0177 | -1.77e-05 | A |  | 0.031 | 0.0177 |  | degradation | ACA @ 50Hz | 0.06 | within limit | within limit
ACA @ 50Hz | 329 | mA | 329.0512 | 329 | mA | 329.019 | -0.0512 | -0.019 | -1.9e-05 | A |  | 0.0512 | 0.019 |  | degradation | ACA @ 50Hz | 0.1974 | within limit | within limit
ACA @ 50Hz | 0.5 | A | 0.50008 | 0.5 | A | 0.50004 | -7.9e-05 | -4.2e-05 | -4.2e-05 | A |  | 7.9e-05 | 4.2e-05 |  | degradation | ACA @ 50Hz | 0.0003 | within limit | within limit
ACA @ 50Hz | 1.09 | A | 1.0902 | 1.09 | A | 1.09 | -0.000194 | -1.7e-05 | -1.7e-05 | A |  | 0.000194 | 1.7e-05 |  | degradation |  | 0.000654 | within limit | within limit
ACA @ 50Hz | 2 | A | 2.0006 | 2 | A | 1.9999 | -0.00058 | 0.00011 | 0.00011 | A |  | 0.00058 | 0.00011 |  | degradation |  | 0.0012 | within limit | within limit
ACA @ 50Hz | 2.9 | A | 2.9006 | 2.9 | A | 2.8996 | -0.000598 | 0.00042 | 0.00042 | A |  | 0.000598 | 0.00042 |  | degradation |  | 0.00174 | within limit | within limit
ACA @ 50Hz | 5 | A | 5.0012 | 5 | A | 4.9995 | -0.001163 | 0.00046 | 0.00046 | A |  | 0.001163 | 0.00046 |  | degradation |  | 0.003 | within limit | within limit
ACA @ 50Hz | 10.9 | A | 10.9019 | 10.9 | A | 10.9046 | -0.001867 | -0.00455 | -0.00455 | A |  | 0.001867 | 0.00455 |  | improved |  | 0.00654 | within limit | within limit
ACA @ 50Hz | 20 | A | 20.0016 | 20 | A | 20.0041 | -0.001563 | -0.00412 | -0.00412 | A |  | 0.001563 | 0.00412 |  | improved |  | 0.012 | within limit | within limit
ACA @ 1kHz | 33 | µA | 33.049 | 33 | µA | 32.9978 | -0.049 | 0.0022 | 2.2e-09 | A |  | 0.049 | 0.0022 |  | degradation | ACA @ 1kHz | 0.0198 | out of limit | within limit
ACA @ 1kHz | 100 | µA | 100.023 | 100 | µA | 100.008 | -0.023 | -0.008 | -8e-09 | A |  | 0.023 | 0.008 |  | degradation | ACA @ 1kHz | 0.06 | within limit | within limit
ACA @ 1kHz | 200 | µA | 200.026 | 200 | µA | 200.068 | -0.026 | -0.068 | -6.8e-08 | A |  | 0.026 | 0.068 |  | degradation | ACA @ 1kHz | 0.12 | within limit | within limit
ACA @ 1kHz | 329 | µA | 329.04 | 329 | µA | 329.071 | -0.04 | -0.071 | -7.1e-08 | A |  | 0.04 | 0.071 |  | degradation | ACA @ 1kHz | 0.1974 | within limit | within limit
ACA @ 1kHz | 1 | mA | 0.99993 | 1 | mA | 1.0002 | 7e-05 | -0.000152 | -1.52e-07 | A |  | 7e-05 | 0.000152 |  | improved |  | 0.0006 | within limit | within limit
ACA @ 1kHz | 3.29 | mA | 3.2913 | 3.29 | mA | 3.29 | -0.0013 | 4e-05 | 4e-08 | A |  | 0.0013 | 4e-05 |  | degradation | ACA @ 1kHz | 0.001974 | within limit | within limit
ACA @ 1kHz | 10 | mA | 10.0048 | 10 | mA | 10.0035 | -0.0048 | -0.0035 | -3.5e-06 | A |  | 0.0048 | 0.0035 |  | degradation | ACA @ 1kHz | 0.006 | within limit | within limit
ACA @ 1kHz | 32.9 | mA | 32.9135 | 32.9 | mA | 32.9062 | -0.01348 | -0.0062 | -6.2e-06 | A |  | 0.01348 | 0.0062 |  | degradation | ACA @ 1kHz | 0.01974 | within limit | within limit
ACA @ 1kHz | 100 | mA | 100.0234 | 100 | mA | 100.0356 | -0.0234 | -0.0356 | -3.56e-05 | A |  | 0.0234 | 0.0356 |  | degradation | ACA @ 1kHz | 0.06 | within limit | within limit
ACA @ 1kHz | 329 | mA | 329.114 | 329 | mA | 329.034 | -0.114 | -0.034 | -3.4e-05 | A |  | 0.114 | 0.034 |  | degradation | ACA @ 1kHz | 0.1974 | within limit | within limit
ACA @ 1kHz | 0.5 | A | 0.50013 | 0.5 | A | 0.50004 | -0.000132 | -3.9e-05 | -3.9e-05 | A |  | 0.000132 | 3.9e-05 |  | degradation | ACA @ 1kHz | 0.0003 | within limit | within limit
ACA @ 1kHz | 1.09 | A | 1.0904 | 1.09 | A | 1.09 | -0.000374 | -1.8e-05 | -1.8e-05 | A |  | 0.000374 | 1.8e-05 |  | degradation | ACA @ 1kHz | 0.000654 | within limit | within limit
ACA @ 1kHz | 2 | A | 2.0006 | 2 | A | 1.9999 | -0.00056 | 8e-05 | 8e-05 | A |  | 0.00056 | 8e-05 |  | degradation | ACA @ 1kHz | 0.0012 | within limit | within limit
ACA @ 1kHz | 2.9 | A | 2.901 | 2.9 | A | 2.9005 | -0.00099 | -0.00046 | -0.00046 | A |  | 0.00099 | 0.00046 |  | degradation | ACA @ 1kHz | 0.00174 | within limit | within limit
ACA @ 1kHz | 5 | A | 5.0018 | 5 | A | 4.9996 | -0.001799 | 0.00038 | 0.00038 | A |  | 0.001799 | 0.00038 |  | degradation | ACA @ 1kHz | 0.003 | within limit | within limit
ACA @ 1kHz |  | A |  | 10 | A | 9.9985 |  | 0.00151 | 0.00151 | A |  |  | 0.00151 |  | improved |  | 0 | out of limit | out of limit
ACA @ 1kHz | 20 | A | 20.004 | 20 | A | 19.9992 | -0.004025 | 0.00082 | 0.00082 | A |  | 0.004025 | 0.00082 |  | degradation | ACA @ 1kHz | 0.012 | within limit | within limit
ACA @ 10kHz | 100 | µA | 99.868 | 100 | µA | 99.9943 | 0.132 | 0.0057 | 5.7e-09 | A |  | 0.132 | 0.0057 |  | improved |  | 0.06 | out of limit | within limit
ACA @ 10kHz | 200 | µA | 199.759 | 200 | µA | 199.924 | 0.241 | 0.076 | 7.6e-08 | A |  | 0.241 | 0.076 |  | improved |  | 0.12 | out of limit | within limit
ACA @ 10kHz | 329 | µA | 328.795 | 329 | µA | 328.911 | 0.205 | 0.089 | 8.9e-08 | A |  | 0.205 | 0.089 |  | improved |  | 0.1974 | out of limit | within limit
ACA @ 10kHz | 1 | mA | 1 | 1 | mA | 0.99978 | -7e-06 | 0.000216 | 2.16e-07 | A |  | 7e-06 | 0.000216 |  | degradation | ACA @ 10kHz | 0.0006 | within limit | within limit
ACA @ 10kHz | 3.29 | mA | 3.2903 | 3.29 | mA | 3.289 | -0.00032 | 0.00098 | 9.8e-07 | A |  | 0.00032 | 0.00098 |  | degradation | ACA @ 10kHz | 0.001974 | within limit | within limit
ACA @ 10kHz | 10 | mA | 10.0203 | 10 | mA | 10.0022 | -0.02034 | -0.00224 | -2.24e-06 | A |  | 0.02034 | 0.00224 |  | degradation | ACA @ 10kHz | 0.006 | out of limit | within limit
ACA @ 10kHz | 32.9 | mA | 32.9553 | 32.9 | mA | 32.9244 | -0.0553 | -0.0244 | -2.44e-05 | A |  | 0.0553 | 0.0244 |  | degradation | ACA @ 10kHz | 0.01974 | out of limit | out of limit
ACA @ 10kHz | 100 | mA | 100.2257 | 100 | mA | 100.0047 | -0.2257 | -0.0047 | -4.7e-06 | A |  | 0.2257 | 0.0047 |  | degradation | ACA @ 10kHz | 0.06 | out of limit | within limit
ACA @ 10kHz | 329 | mA | 329.702 | 329 | mA | 329.098 | -0.702 | -0.098 | -9.8e-05 | A |  | 0.702 | 0.098 |  | degradation | ACA @ 10kHz | 0.1974 | out of limit | within limit
RESISTANCE(4 wire) | 1 | Ω | 1 | 1 | Ω | 0.99996 | 0 | 3.95e-05 | 3.95e-05 | Ω |  | 0 | 3.95e-05 |  | degradation | RESISTANCE(4 wire) | 2.8e-05 | within limit | out of limit
RESISTANCE(4 wire) | 5 | Ω | 5 | 5 | Ω | 4.9998 | -2e-05 | 0.000158 | 0.000158 | Ω |  | 2e-05 | 0.000158 |  | degradation | RESISTANCE(4 wire) | 0.00014 | within limit | out of limit
RESISTANCE(4 wire) | 10.9 | Ω | 10.9001 | 10.9 | Ω | 10.8998 | -0.00013 | 0.000202 | 0.000202 | Ω |  | 0.00013 | 0.000202 |  | degradation | RESISTANCE(4 wire) | 0.0003052 | within limit | within limit
RESISTANCE(4 wire) | 20 | Ω | 20.0002 | 20 | Ω | 19.9997 | -0.0002 | 0.00026 | 0.00026 | Ω |  | 0.0002 | 0.00026 |  | degradation | RESISTANCE(4 wire) | 0.00056 | within limit | within limit
RESISTANCE(4 wire) | 32.9 | Ω | 32.9002 | 32.9 | Ω | 32.8994 | -0.00022 | 0.00056 | 0.00056 | Ω |  | 0.00022 | 0.00056 |  | degradation | RESISTANCE(4 wire) | 0.0009212 | within limit | within limit
RESISTANCE(4 wire) | 50 | Ω | 50 | 50 | Ω | 49.9994 | -3e-05 | 0.00059 | 0.00059 | Ω |  | 3e-05 | 0.00059 |  | degradation | RESISTANCE(4 wire) | 0.0014 | within limit | within limit
RESISTANCE(4 wire) | 109.9 | Ω | 109.9 | 109.9 | Ω | 109.8987 | 5e-05 | 0.00127 | 0.00127 | Ω |  | 5e-05 | 0.00127 |  | degradation | RESISTANCE(4 wire) | 0.0030772 | within limit | within limit
RESISTANCE(4 wire) | 200 | Ω | 200.0029 | 200 | Ω | 199.9975 | -0.00288 | 0.0025 | 0.0025 | Ω |  | 0.00288 | 0.0025 |  | degradation | RESISTANCE(4 wire) | 0.0056 | within limit | within limit
RESISTANCE(4 wire) | 329.9 | Ω | 329.9035 | 329.9 | Ω | 329.894 | -0.00353 | 0.006 | 0.006 | Ω |  | 0.00353 | 0.006 |  | degradation | RESISTANCE(4 wire) | 0.0092372 | within limit | within limit
RESISTANCE(4 wire) | 0.5 | kΩ | 0.5 | 0.5 | kΩ | 0.49999 | -1.2e-06 | 5.3e-06 | 0.0053 | Ω |  | 1.2e-06 | 5.3e-06 |  | degradation | RESISTANCE(4 wire) | 1.4e-05 | within limit | within limit
RESISTANCE(4 wire) | 1.09 | kΩ | 1.09 | 1.09 | kΩ | 1.09 | -1.9e-06 | 1.4e-05 | 0.014 | Ω |  | 1.9e-06 | 1.4e-05 |  | degradation | RESISTANCE(4 wire) | 3.052e-05 | within limit | within limit
RESISTANCE(4 wire) | 3.29 | kΩ | 3.29 | 3.29 | kΩ | 3.2899 | 1.55e-05 | 7e-05 | 0.07 | Ω |  | 1.55e-05 | 7e-05 |  | degradation | RESISTANCE(4 wire) | 9.212e-05 | within limit | within limit
RESISTANCE(4 wire) | 5 | kΩ | 5 | 5 | kΩ | 4.9999 | 3.2e-05 | 8.5e-05 | 0.085 | Ω |  | 3.2e-05 | 8.5e-05 |  | degradation | RESISTANCE(4 wire) | 0.00014 | within limit | within limit
RESISTANCE(4 wire) | 10.9 | kΩ | 10.8999 | 10.9 | kΩ | 10.8998 | 6.5e-05 | 0.000189 | 0.189 | Ω |  | 6.5e-05 | 0.000189 |  | degradation | RESISTANCE(4 wire) | 0.0003052 | within limit | within limit
RESISTANCE(4 wire) | 20 | kΩ | 19.9999 | 20 | kΩ | 19.9997 | 5.4e-05 | 0.00027 | 0.27 | Ω |  | 5.4e-05 | 0.00027 |  | degradation | RESISTANCE(4 wire) | 0.00056 | within limit | within limit
RESISTANCE(4 wire) | 32.9 | kΩ | 32.8999 | 32.9 | kΩ | 32.9005 | 0.000129 | -0.00046 | -0.46 | Ω |  | 0.000129 | 0.00046 |  | improved |  | 0.0009212 | within limit | within limit
RESISTANCE(4 wire) | 50 | kΩ | 49.9998 | 50 | kΩ | 49.9992 | 0.00021 | 0.00079 | 0.79 | Ω |  | 0.00021 | 0.00079 |  | degradation | RESISTANCE(4 wire) | 0.0014 | within limit | within limit
RESISTANCE(2 wire) | 109.9 | kΩ | 109.8997 | 109.9 | kΩ | 109.8986 | 0.00033 | 0.00143 | 1.43 | Ω |  | 0.00033 | 0.00143 |  | degradation | RESISTANCE(2 wire) | 0.0030772 | within limit | within limit
RESISTANCE(2 wire) | 200 | kΩ | 199.9994 | 200 | kΩ | 199.9966 | 0.00056 | 0.0034 | 3.4 | Ω |  | 0.00056 | 0.0034 |  | degradation | RESISTANCE(2 wire) | 0.0056 | within limit | within limit
RESISTANCE(2 wire) | 329.9 | kΩ | 329.8989 | 329.9 | kΩ | 329.896 | 0.00108 | 0.004 | 4 | Ω |  | 0.00108 | 0.004 |  | degradation | RESISTANCE(2 wire) | 0.0092372 | within limit | within limit
RESISTANCE(2 wire) | 0.5 | MΩ | 0.5 | 0.5 | MΩ | 0.49999 | 2.3e-06 | 6.1e-06 | 6.1 | Ω |  | 2.3e-06 | 6.1e-06 |  | degradation | RESISTANCE(2 wire) | 1.4e-05 | within limit | within limit
RESISTANCE(2 wire) | 1.09 | MΩ | 1.09 | 1.09 | MΩ | 1.09 | 5e-06 | 1.08e-05 | 10.8 | Ω |  | 5e-06 | 1.08e-05 |  | degradation | RESISTANCE(2 wire) | 3.052e-05 | within limit | within limit
RESISTANCE(2 wire) | 2 | MΩ | 2 | 2 | MΩ | 2 | 4.06e-05 | 4e-05 | 40 | Ω |  | 4.06e-05 | 4e-05 |  | degradation | RESISTANCE(2 wire) | 5.6e-05 | within limit | within limit
RESISTANCE(2 wire) | 3.29 | MΩ | 3.29 | 3.29 | MΩ | 3.29 | 3.38e-05 | 4.7e-05 | 47 | Ω |  | 3.38e-05 | 4.7e-05 |  | degradation | RESISTANCE(2 wire) | 9.212e-05 | within limit | within limit
RESISTANCE(2 wire) | 5 | MΩ | 4.9999 | 5 | MΩ | 4.9999 | 0.000126 | 0.000137 | 137 | Ω |  | 0.000126 | 0.000137 |  | degradation | RESISTANCE(2 wire) | 0.00014 | within limit | within limit
RESISTANCE(2 wire) | 10.9 | MΩ | 10.8996 | 10.9 | MΩ | 10.8999 | 0.000418 | 0.000116 | 116 | Ω |  | 0.000418 | 0.000116 |  | degradation | RESISTANCE(2 wire) | 0.0003052 | out of limit | within limit
RESISTANCE(2 wire) | 20 | MΩ | 19.9962 | 20 | MΩ | 19.9996 | 0.003814 | 0.00037 | 370 | Ω |  | 0.003814 | 0.00037 |  | degradation | RESISTANCE(2 wire) | 0.00056 | out of limit | within limit
RESISTANCE(2 wire) | 32.9 | MΩ | 32.8974 | 32.9 | MΩ | 32.9003 | 0.002639 | -0.00031 | -310 | Ω |  | 0.002639 | 0.00031 |  | improved |  | 0.0009212 | out of limit | within limit
RESISTANCE(2 wire) | 50 | MΩ | 49.9941 | 50 | MΩ | 49.9954 | 0.00585 | 0.00456 | 4560 | Ω |  | 0.00585 | 0.00456 |  | degradation | RESISTANCE(2 wire) | 0.0014 | out of limit | out of limit
RESISTANCE(2 wire) | 109.9 | MΩ | 109.8622 | 109.9 | MΩ | 109.9022 | 0.03784 | -0.00216 | -2160 | Ω |  | 0.03784 | 0.00216 |  | improved |  | 0.0030772 | out of limit | within limit
RESISTANCE(2 wire) | 200 | MΩ | 199.8241 | 200 | MΩ | 199.9992 | 0.1759 | 0.0008 | 800 | Ω |  | 0.1759 | 0.0008 |  | degradation | RESISTANCE(2 wire) | 0.0056 | out of limit | within limit
RESISTANCE(2 wire) | 329.9 | MΩ | 329.6876 | 329.9 | MΩ | 330.0379 | 0.2124 | -0.1379 | -137900 | Ω |  | 0.2124 | 0.1379 |  | improved |  | 0.0092372 | out of limit | out of limit
RESISTANCE(2 wire) | 1000 | MΩ | 999.143 | 1000 | MΩ | 1000.0895 | 0.857 | -0.0895 | -89500 | Ω |  | 0.857 | 0.0895 |  | improved |  | 0.028 | out of limit | out of limit
FREQUENCY |  |  |  | 45 | Hz | 45 |  | 1.2e-05 | 1.2e-05 | Hz |  |  | 1.2e-05 |  | improved |  | 0 | out of limit | out of limit
FREQUENCY | 50 | Hz | 50 | 50 | Hz | 50 | 2.46e-05 | 1.6e-05 | 1.6e-05 | Hz |  | 2.46e-05 | 1.6e-05 |  | improved |  | 0.00125 | within limit | within limit
FREQUENCY | 100 | Hz | 100 | 100 | Hz | 100 | 4.2e-05 | 1e-05 | 1e-05 | Hz |  | 4.2e-05 | 1e-05 |  | improved |  | 0.0025 | within limit | within limit
FREQUENCY | 500 | Hz | 499.9998 | 500 | Hz | 499.9998 | 0.0002 | 0.00018 | 0.00018 | Hz |  | 0.0002 | 0.00018 |  | improved |  | 0.0125 | within limit | within limit
FREQUENCY | 1100 | Hz | 1099.9995 | 1100 | Hz | 1099.9996 | 0.0005 | 0.0004 | 0.0004 | Hz |  | 0.0005 | 0.0004 |  | improved |  | 0.0275 | within limit | within limit
FREQUENCY |  |  |  | 0.5 | kHz | 0.5 |  | 1.9e-07 | 0.00019 | Hz |  |  | 1.9e-07 |  | improved |  | 0 | out of limit | out of limit
FREQUENCY | 11.9 | kHz | 11.9 | 11.9 | kHz | 11.9 | 6e-06 | 5e-06 | 0.005 | Hz |  | 6e-06 | 5e-06 |  | degradation | FREQUENCY | 0.0002975 | within limit | within limit
FREQUENCY | 50 | kHz | 50 | 50 | kHz | 50 | 2e-05 | 2.2e-05 | 0.022 | Hz |  | 2e-05 | 2.2e-05 |  | degradation | FREQUENCY | 0.00125 | within limit | within limit
FREQUENCY | 119 | kHz | 118.9999 | 119 | kHz | 118.9999 | 6e-05 | 0.00014 | 0.14 | Hz |  | 6e-05 | 0.00014 |  | degradation | FREQUENCY | 0.002975 | within limit | within limit
FREQUENCY | 500 | kHz | 499.9998 | 500 | kHz | 499.9998 | 0.00024 | 0.00021 | 0.21 | Hz |  | 0.00024 | 0.00021 |  | degradation | FREQUENCY | 0.0125 | within limit | within limit
FREQUENCY | 1 | MHz | 1 | 1 | MHz | 1 | 5e-07 | -2.1e-06 | -2.1 | Hz |  | 5e-07 | 2.1e-06 |  | improved |  | 2.5e-05 | within limit | within limit
TEMPERATURE RTD | -199 | °C | -198.9944 | -200 | °C | -199.9981 | -0.0056 | -0.0019 | -0.0019 | °C |  | 0.0056 | 0.0019 |  | degradation | TEMPERATURE RTD | 0.03 | within limit | within limit
TEMPERATURE RTD |  |  |  | -50 | °C | -50.0021 |  | 0.0021 | 0.0021 | °C |  |  | 0.0021 |  | improved |  | 0.03 | out of limit | within limit
TEMPERATURE RTD | 0 | °C | -0.0051 | 0 | °C | -0.0035 | 0.0051 | 0.0035 | 0.0035 | °C |  | 0.0051 | 0.0035 |  | improved |  | 0.03 | within limit | within limit
TEMPERATURE RTD |  |  |  | 50 | °C | 49.9958 |  | 0.0042 | 0.0042 | °C |  |  | 0.0042 |  | improved |  | 0.03 | out of limit | within limit
TEMPERATURE RTD |  |  |  | 200 | °C | 199.9915 |  | 0.0085 | 0.0085 | °C |  |  | 0.0085 |  | improved |  | 0.03 | out of limit | within limit
TEMPERATURE RTD | 400 | °C | 399.9776 | 400 | °C | 399.9832 | 0.0224 | 0.0168 | 0.0168 | °C |  | 0.0224 | 0.0168 |  | improved |  | 0.03 | within limit | within limit
TEMPERATURE RTD | 800 | °C | 799.9697 | 800 | °C | 799.9862 | 0.0303 | 0.0138 | 0.0138 | °C |  | 0.0303 | 0.0138 |  | improved |  | 0.03 | out of limit | within limit
CAPACITANCE |  |  |  | 0.5 | nF | 0.44983 |  | 0.050173 | 0 | F |  |  | 0.050173 |  | improved |  | 0 | out of limit | out of limit
CAPACITANCE | 1 | nF | 1.0064 | 1 | nF | 1.0008 | -0.00638 | -0.000756 | 0 | F |  | 0.00638 | 0.000756 |  | degradation | CAPACITANCE | 0.0025 | out of limit | within limit
CAPACITANCE |  |  |  | 1.5 | nF | 1.4991 |  | 0.000914 | 0 | F |  |  | 0.000914 |  | improved |  | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 3.2 | nF | 3.1991 |  | 0.000927 | 0 | F |  |  | 0.000927 |  | improved |  | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 5 | nF | 4.9995 |  | 0.000513 | 0 | F |  |  | 0.000513 |  | improved |  | 0 | out of limit | out of limit
CAPACITANCE | 10 | nF | 10.0617 | 10.9 | nF | 10.9016 | -0.06168 | -0.00164 | 0 | F |  | 0.06168 | 0.00164 |  | degradation | CAPACITANCE | 0.025 | out of limit | within limit
CAPACITANCE |  |  |  | 15 | nF | 15.0032 |  | -0.00325 | 0 | F |  |  | 0.00325 |  | improved |  | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 32.9 | nF | 32.8787 |  | 0.0213 | 0 | F |  |  | 0.0213 |  | improved |  | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 50 | nF | 49.9983 |  | 0.00166 | 0 | F |  |  | 0.00166 |  | improved |  | 0 | out of limit | out of limit
CAPACITANCE | 100 | nF | 97.7023 | 100.9 | nF | 100.8915 | 2.2977 | 0.0085 | 0 | F |  | 2.2977 | 0.0085 |  | improved |  | 0.25 | out of limit | within limit
CAPACITANCE |  |  |  | 150 | nF | 150.0546 |  | -0.0546 | 0 | F |  |  | 0.0546 |  | improved |  | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 329.9 | nF | 329.8638 |  | 0.0362 | 0 | F |  |  | 0.0362 |  | improved |  | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 0.5 | µF | 0.50227 |  | -0.0022748 | #N/A | #N/A |  |  | 0.0022748 |  | #N/A | #N/A | 0 | out of limit | out of limit
CAPACITANCE | 1 | µF | 1.0023 | 1.09 | µF | 1.0904 | -0.0023 | -0.000359 | #N/A | #N/A |  | 0.0023 | 0.000359 |  | #N/A | #N/A | 0.0025 | within limit | within limit
CAPACITANCE |  |  |  | 1.5 | µF | 1.4995 |  | 0.000519 | #N/A | #N/A |  |  | 0.000519 |  | #N/A | #N/A | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 3.29 | µF | 3.2887 |  | 0.00126 | #N/A | #N/A |  |  | 0.00126 |  | #N/A | #N/A | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 5 | µF | 4.9971 |  | 0.00289 | #N/A | #N/A |  |  | 0.00289 |  | #N/A | #N/A | 0 | out of limit | out of limit
CAPACITANCE | 10 | µF | 9.55 | 10.9 | µF | 10.8846 | 0.45 | 0.01538 | #N/A | #N/A |  | 0.45 | 0.01538 |  | #N/A | #N/A | 0.025 | out of limit | within limit
CAPACITANCE |  |  |  | 15 | µF | 15.0033 |  | -0.00326 | #N/A | #N/A |  |  | 0.00326 |  | #N/A | #N/A | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 32.9 | µF | 32.8985 |  | 0.00149 | #N/A | #N/A |  |  | 0.00149 |  | #N/A | #N/A | 0 | out of limit | out of limit
CAPACITANCE |  |  |  | 50 | µF | 50.0063 |  | -0.00633 | #N/A | #N/A |  |  | 0.00633 |  | #N/A | #N/A | 0 | out of limit | out of limit
CAPACITANCE | 100 | µF | 95.49 | 109.9 | µF | 109.8992 | 4.51 | 0.0008 | #N/A | #N/A |  | 4.51 | 0.0008 |  | #N/A | #N/A | 0.25 | out of limit | within limit
T/C J-TYPE(Source) | -200 | °C | -200.023 | -200 | °C | -200.12 | 0.023 | 0.12 | 0.12 | °C |  | 0.023 | 0.12 |  | degradation | T/C J-TYPE(Source) | 0.14 | within limit | within limit
T/C J-TYPE(Source) |  |  |  | 0 | °C | -0.02 |  | 0.02 | 0.02 | °C |  |  | 0.02 |  | improved |  | 0.14 | out of limit | within limit
T/C J-TYPE(Source) |  |  |  | 100 | °C | 99.89 |  | 0.11 | 0.11 | °C |  |  | 0.11 |  | improved |  | 0.14 | out of limit | within limit
T/C J-TYPE(Source) |  |  |  | 600 | °C | 599.85 |  | 0.15 | 0.15 | °C |  |  | 0.15 |  | improved |  | 0.14 | out of limit | out of limit
T/C J-TYPE(Source) |  |  |  | 1050 | °C | 1049.85 |  | 0.15 | 0.15 | °C |  |  | 0.15 |  | improved |  | 0.14 | out of limit | out of limit
T/C K-TYPE(Source) | -200 | °C | -200.006 | -200 | °C | -199.99 | 0.006 | -0.01 | -0.01 | °C |  | 0.006 | 0.01 |  | improved |  | 0.14 | within limit | within limit
T/C K-TYPE(Source) |  |  |  | 0 | °C | 0.09 |  | -0.09 | -0.09 | °C |  |  | 0.09 |  | improved |  | 0.14 | out of limit | within limit
T/C K-TYPE(Source) |  |  |  | 100 | °C | 100.07 |  | -0.07 | -0.07 | °C |  |  | 0.07 |  | improved |  | 0.14 | out of limit | within limit
T/C K-TYPE(Source) | 500 | °C | 499.99 | 500 | °C | 500.07 | 0.01 | -0.07 | -0.07 | °C |  | 0.01 | 0.07 |  | improved |  | 0.14 | within limit | within limit
T/C K-TYPE(Source) | 1000 | °C | 999.962 | 1000 | °C | 1000.09 | 0.038 | -0.09 | -0.09 | °C |  | 0.038 | 0.09 |  | improved |  | 0.14 | within limit | within limit
T/C K-TYPE(Source) |  |  |  | 1300 | °C | 1300.14 |  | -0.14 | -0.14 | °C |  |  | 0.14 |  | improved |  | 0.14 | out of limit | out of limit
T/C R-TYPE(Source) |  |  |  | 0 | °C | 0.15 |  | -0.15 | -0.15 | °C |  |  | 0.15 |  | improved |  | 0.14 | out of limit | out of limit
T/C R-TYPE(Source) | 100 | °C | 100.013 | 100 | °C | 100.15 | -0.013 | -0.15 | -0.15 | °C |  | 0.013 | 0.15 |  | improved |  | 0.14 | within limit | out of limit
T/C R-TYPE(Source) | 500 | °C | 500 | 500 | °C | 500.17 | 0 | -0.17 | -0.17 | °C |  | 0 | 0.17 |  | improved |  | 0.14 | within limit | out of limit
T/C R-TYPE(Source) | 1000 | °C | 999.977 | 1000 | °C | 1000.16 | 0.023 | -0.16 | -0.16 | °C |  | 0.023 | 0.16 |  | improved |  | 0.14 | within limit | out of limit
T/C R-TYPE(Source) |  |  |  | 1750 | °C | 1750.18 |  | -0.18 | -0.18 | °C |  |  | 0.18 |  | improved |  | 0.14 | out of limit | out of limit
T/C S-TYPE(Source) |  |  |  | 0 | °C | 0.16 |  | -0.16 | -0.16 | °C |  |  | 0.16 |  | improved |  | 0.14 | out of limit | out of limit
T/C S-TYPE(Source) | 100 | °C | 99.957 | 100 | °C | 100.12 | 0.043 | -0.12 | -0.12 | °C |  | 0.043 | 0.12 |  | improved |  | 0.14 | within limit | within limit
T/C S-TYPE(Source) | 500 | °C | 499.95 | 500 | °C | 500.16 | 0.05 | -0.16 | -0.16 | °C |  | 0.05 | 0.16 |  | improved |  | 0.14 | within limit | out of limit
T/C S-TYPE(Source) | 1000 | °C | 999.973 | 1000 | °C | 1000.15 | 0.027 | -0.15 | -0.15 | °C |  | 0.027 | 0.15 |  | improved |  | 0.14 | within limit | out of limit
T/C S-TYPE(Source) |  |  |  | 1750 | °C | 1750.18 |  | -0.18 | -0.18 | °C |  |  | 0.18 |  | improved |  | 0.14 | out of limit | out of limit
T/C T-TYPE(Source) |  |  |  | -200 | °C | -200.1 |  | 0.1 | 0.1 | °C |  |  | 0.1 |  | improved |  | 0.14 | out of limit | within limit
T/C T-TYPE(Source) |  |  |  | 0 | °C | -0.02 |  | 0.02 | 0.02 | °C |  |  | 0.02 |  | improved |  | 0.14 | out of limit | within limit
T/C T-TYPE(Source) | 100 | °C | 99.989 | 100 | °C | 100 | 0.011 | 0 | 0 | °C |  | 0.011 | 0 |  | improved |  | 0.14 | within limit | within limit
T/C T-TYPE(Source) |  |  |  | 400 | °C | 399.96 |  | 0.04 | 0.04 | °C |  |  | 0.04 |  | improved |  | 0.14 | out of limit | within limit
T/C N-TYPE(Source) | -200 | °C | -200.03 | -200 | °C | -199.9 | 0.03 | -0.1 | -0.1 | °C |  | 0.03 | 0.1 |  | improved |  | 0.14 | within limit | within limit
T/C N-TYPE(Source) |  |  |  | 0 | °C | 0.12 |  | -0.12 | -0.12 | °C |  |  | 0.12 |  | improved |  | 0.14 | out of limit | within limit
T/C N-TYPE(Source) | 100 | °C | 100.027 | 100 | °C | 100.17 | -0.027 | -0.17 | -0.17 | °C |  | 0.027 | 0.17 |  | improved |  | 0.14 | within limit | out of limit
T/C N-TYPE(Source) | 500 | °C | 499.987 | 500 | °C | 500.16 | 0.013 | -0.16 | -0.16 | °C |  | 0.013 | 0.16 |  | improved |  | 0.14 | within limit | out of limit
T/C N-TYPE(Source) | 1000 | °C | 999.964 | 1000 | °C | 1000.16 | 0.036 | -0.16 | -0.16 | °C |  | 0.036 | 0.16 |  | improved |  | 0.14 | within limit | out of limit
T/C N-TYPE(Source) |  |  |  | 1300 | °C | 1300.19 |  | -0.19 | -0.19 | °C |  |  | 0.19 |  | improved |  | 0.14 | out of limit | out of limit
T/C B-TYPE(Measure) |  |  |  | 600.11 | °C | 600 |  | 0.11 | 0.11 | °C |  |  | 0.11 |  | improved |  | 0.14 | out of limit | within limit
T/C B-TYPE(Measure) |  |  |  | 1200.15 | °C | 1200 |  | 0.15 | 0.15 | °C |  |  | 0.15 |  | improved |  | 0.14 | out of limit | out of limit
T/C B-TYPE(Measure) |  |  |  | 1800.08 | °C | 1800 |  | 0.08 | 0.08 | °C |  |  | 0.08 |  | improved |  | 0.14 | out of limit | within limit
T/C E-TYPE(Measure) |  |  |  | -199.96 | °C | -200 |  | 0.04 | 0.04 | °C |  |  | 0.04 |  | improved |  | 0.14 | out of limit | within limit
T/C E-TYPE(Measure) |  |  |  | 0.03 | °C | 0 |  | 0.03 | 0.03 | °C |  |  | 0.03 |  | improved |  | 0.14 | out of limit | within limit
T/C E-TYPE(Measure) |  |  |  | 100.07 | °C | 100 |  | 0.07 | 0.07 | °C |  |  | 0.07 |  | improved |  | 0.14 | out of limit | within limit
T/C E-TYPE(Measure) |  |  |  | 500.09 | °C | 500 |  | 0.09 | 0.09 | °C |  |  | 0.09 |  | improved |  | 0.14 | out of limit | within limit
T/C E-TYPE(Measure) |  |  |  | 1000.14 | °C | 1000 |  | 0.14 | 0.14 | °C |  |  | 0.14 |  | improved |  | 0.14 | out of limit | within limit
T/C J-TYPE(Measure) | -199.94 | °C | -200 | -199.86 | °C | -200 | 0.06 | 0.14 | 0.14 | °C |  | 0.06 | 0.14 |  | degradation | T/C J-TYPE(Measure) | 0.14 | within limit | within limit
T/C J-TYPE(Measure) |  |  |  | 0.02 | °C | 0 |  | 0.02 | 0.02 | °C |  |  | 0.02 |  | improved |  | 0.14 | out of limit | within limit
T/C J-TYPE(Measure) |  |  |  | 100.08 | °C | 100 |  | 0.08 | 0.08 | °C |  |  | 0.08 |  | improved |  | 0.14 | out of limit | within limit
T/C J-TYPE(Measure) | 600 | °C | 600 | 600.1 | °C | 600 | 0 | 0.1 | 0.1 | °C |  | 0 | 0.1 |  | degradation | T/C J-TYPE(Measure) | 0.14 | within limit | within limit
T/C J-TYPE(Measure) |  |  |  | 1050.1 | °C | 1050 |  | 0.1 | 0.1 | °C |  |  | 0.1 |  | improved |  | 0.14 | out of limit | within limit
T/C K-TYPE(Measure) | -199.89 | °C | -200 | -199.87 | °C | -200 | -0.02 | 0.13 | 0.13 | °C |  | 0.02 | #REF! | #REF! | degradation | T/C K-TYPE(Measure) | 0.14 | within limit | #REF!
T/C K-TYPE(Measure) |  |  |  | 0.12 | °C | 0 | -0.12 | 0.12 | 0.12 | °C |  | 0.12 | #REF! | #REF! | degradation | T/C K-TYPE(Measure) | 0.14 | within limit | #REF!
T/C K-TYPE(Measure) |  |  |  | 100.16 | °C | 100 | -100.16 | 0.16 | 0.16 | °C |  | 100.16 | #REF! | #REF! | degradation | T/C K-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C K-TYPE(Measure) |  |  |  | 500.13 | °C | 500 | -500.13 | 0.13 | 0.13 | °C |  | 500.13 | #REF! | #REF! | degradation | T/C K-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C K-TYPE(Measure) |  |  |  | 1000.15 | °C | 1000 | -1000.15 | 0.15 | 0.15 | °C |  | 1000.15 | #REF! | #REF! | degradation | T/C K-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C K-TYPE(Measure) |  |  |  | 1300.19 | °C | 1300 | -1300.19 | 0.19 | 0.19 | °C |  | 1300.19 | #REF! | #REF! | degradation | T/C K-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C R-TYPE(Measure) |  |  |  | 0.08 | °C | 0 | -0.08 | 0.08 | 0.08 | °C |  | 0.08 | #REF! | #REF! | degradation | T/C R-TYPE(Measure) | 0.14 | within limit | #REF!
T/C R-TYPE(Measure) |  |  |  | 99.88 | °C | 100 | -99.88 | -0.12 | -0.12 | °C |  | 99.88 | #REF! | #REF! | degradation | T/C R-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C R-TYPE(Measure) |  |  |  | 499.9 | °C | 500 | -499.9 | -0.1 | -0.1 | °C |  | 499.9 | #REF! | #REF! | degradation | T/C R-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C R-TYPE(Measure) |  |  |  | 999.89 | °C | 1000 | -999.89 | -0.11 | -0.11 | °C |  | 999.89 | #REF! | #REF! | degradation | T/C R-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C R-TYPE(Measure) |  |  |  | 1749.88 | °C | 1750 | -1749.88 | -0.12 | -0.12 | °C |  | 1749.88 | #REF! | #REF! | degradation | T/C R-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C S-TYPE(Measure) |  |  |  | 0.02 | °C | 0 | -0.02 | 0.02 | 0.02 | °C |  | 0.02 | #REF! | #REF! | degradation | T/C S-TYPE(Measure) | 0.14 | within limit | #REF!
T/C S-TYPE(Measure) |  |  |  | 99.91 | °C | 100 | -99.91 | -0.09 | -0.09 | °C |  | 99.91 | #REF! | #REF! | degradation | T/C S-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C S-TYPE(Measure) |  |  |  | 499.9 | °C | 500 | -499.9 | -0.1 | -0.1 | °C |  | 499.9 | #REF! | #REF! | degradation | T/C S-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C S-TYPE(Measure) |  |  |  | 999.93 | °C | 1000 | -999.93 | -0.07 | -0.07 | °C |  | 999.93 | #REF! | #REF! | degradation | T/C S-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C S-TYPE(Measure) |  |  |  | 1749.89 | °C | 1750 | -1749.89 | -0.11 | -0.11 | °C |  | 1749.89 | #REF! | #REF! | degradation | T/C S-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C T-TYPE(Measure) |  |  |  | -199.94 | °C | -200 | 199.94 | 0.06 | 0.06 | °C |  | 199.94 | #REF! | #REF! | improved |  | 0.14 | out of limit | #REF!
T/C T-TYPE(Measure) |  |  |  | 0.15 | °C | 0 | -0.15 | 0.15 | 0.15 | °C |  | 0.15 | #REF! | #REF! | degradation | T/C T-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C T-TYPE(Measure) | 100.04 | °C | 100 | 100.13 | °C | 100 | -0.09 | 0.13 | 0.13 | °C |  | 0.09 | #REF! | #REF! | degradation | T/C T-TYPE(Measure) | 0.14 | within limit | #REF!
T/C T-TYPE(Measure) | 399.99 | °C | 400 | 400.08 | °C | 400 | -0.09 | 0.08 | 0.08 | °C |  | 0.09 | #REF! | #REF! | degradation | T/C T-TYPE(Measure) | 0.14 | within limit | #REF!
T/C N-TYPE(Measure) | 199.87 | °C | -200 | -200.06 | °C | -200 | 399.93 | -0.06 | -0.06 | °C |  | 399.93 | #REF! | #REF! | improved |  | 0.14 | out of limit | #REF!
T/C N-TYPE(Measure) |  |  |  | -0.08 | °C | 0 | 0.08 | -0.08 | -0.08 | °C |  | 0.08 | #REF! | #REF! | improved |  | 0.14 | within limit | #REF!
T/C N-TYPE(Measure) |  |  |  | 99.86 | °C | 100 | -99.86 | -0.14 | -0.14 | °C |  | 99.86 | #REF! | #REF! | degradation | T/C N-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C N-TYPE(Measure) |  |  |  | 499.87 | °C | 500 | -499.87 | -0.13 | -0.13 | °C |  | 499.87 | #REF! | #REF! | degradation | T/C N-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C N-TYPE(Measure) |  |  |  | 999.88 | °C | 1000 | -999.88 | -0.12 | -0.12 | °C |  | 999.88 | #REF! | #REF! | degradation | T/C N-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C N-TYPE(Measure) |  |  |  | 1299.93 | °C | 1300 | -1299.93 | -0.07 | -0.07 | °C |  | 1299.93 | #REF! | #REF! | degradation | T/C N-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C B-TYPE(Measure) | 600.19 | °C | 600 | 599.95 | °C | 600 | 0.24 | -0.05 | -0.05 | °C |  | 0.24 | #REF! | #REF! | improved |  | 0.14 | out of limit | #REF!
T/C B-TYPE(Measure) | 1200.08 | °C | 1200 | 1199.91 | °C | 1200 | 0.17 | -0.09 | -0.09 | °C |  | 0.17 | #REF! | #REF! | improved |  | 0.14 | out of limit | #REF!
T/C B-TYPE(Measure) | 1800.09 | °C | 1800 | 1799.96 | °C | 1800 | 0.13 | -0.04 | -0.04 | °C |  | 0.13 | #REF! | #REF! | improved |  | 0.14 | within limit | #REF!
T/C E-TYPE(Measure) | -249.84 | °C | -250 | -249.9 | °C | -250 | 0.06 | 0.1 | 0.1 | °C |  | 0.06 | #REF! | #REF! | degradation | T/C E-TYPE(Measure) | 0.14 | within limit | #REF!
T/C E-TYPE(Measure) |  |  |  | 0.07 | °C | 0 | -0.07 | 0.07 | 0.07 | °C |  | 0.07 | #REF! | #REF! | degradation | T/C E-TYPE(Measure) | 0.14 | within limit | #REF!
T/C E-TYPE(Measure) |  |  |  | 100.13 | °C | 100 | -100.13 | 0.13 | 0.13 | °C |  | 100.13 | #REF! | #REF! | degradation | T/C E-TYPE(Measure) | 0.14 | out of limit | #REF!
T/C E-TYPE(Measure) | 500.01 | °C | 500 | 500.15 | °C | 500 | -0.14 | 0.15 | 0.15 | °C |  | 0.14 | #REF! | #REF! | degradation | T/C E-TYPE(Measure) | 0.14 | within limit | #REF!
T/C E-TYPE(Measure) | 1000.05 | °C | 1000 | 1000.01 | °C | 1000 | 0.04 | 0.01 | 0.01 | °C |  | 0.04 | #REF! | #REF! | improved |  | 0.14 | within limit | #REF!
POWER FACTOR(Lead) | 0.2 | PF | 0.2003 | 0.2 | PF | 0.20005 | -0.0003 | -5e-05 | 0 | F |  | 0.0003 | 5e-05 |  | degradation | POWER FACTOR(Lead) |  | out of limit | out of limit
POWER FACTOR(Lead) | 0.5 | PF | 0.5002 | 0.5 | PF | 0.49988 | -0.0002 | 0.00012 | 0 | F |  | 0.0002 | 0.00012 |  | degradation | POWER FACTOR(Lead) |  | out of limit | out of limit
POWER FACTOR(Lead) | 0.8 | PF | 0.8 | 0.8 | PF | 0.79992 | 0 | 8e-05 | 0 | F |  | 0 | 8e-05 |  | degradation | POWER FACTOR(Lead) |  | within limit | out of limit
POWER FACTOR(Lead) | 1 | PF | 1 | 1 | PF | 1 | 0 | 0 | 0 | F |  | 0 | 0 |  | same condition |  |  | within limit | within limit
POWER FACTOR(Lag) | 1 | PF | 1 | 1 | PF | 1 | 0 | 0 | 0 | F |  | 0 | 0 |  | same condition |  |  | within limit | within limit
POWER FACTOR(Lag) | 0.5 | PF | 0.4999 | 0.5 | PF | 0.50021 | 0.0001 | -0.00021 | 0 | F |  | 0.0001 | 0.00021 |  | improved |  |  | out of limit | out of limit
POWER FACTOR(Lag) | 0.8 | PF | 0.7999 | 0.8 | PF | 0.80016 | 0.0001 | -0.00016 | 0 | F |  | 0.0001 | 0.00016 |  | improved |  |  | out of limit | out of limit
POWER FACTOR(Lag) | 0.2 | PF | 0.1999 | 0.2 | PF | 0.20042 | 0.0001 | -0.00042 | 0 | F |  | 0.0001 | 0.00042 |  | improved |  |  | out of limit | out of limit
AC POWER @ UPF |  |  |  | 1 | W | 1.0008 |  | -0.0008 | -8e-05 | W |  |  | 0.0008 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ UPF |  |  |  | 200 | W | 199.892 |  | 0.108 | 0.0108 | W |  |  | 0.108 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ UPF |  |  |  | 15 | W | 14.992 |  | 0.008 | 0.0008 | W |  |  | 0.008 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ UPF | 150 | W | 150.02 | 150 | W | 150.048 | -0.02 | -0.048 | -0.0048 | W |  | 0.02 | 0.048 |  | degradation | AC POWER @ UPF | 0.12 | within limit | within limit
AC POWER @ UPF | 750 | W | 750.14 | 750 | W | 749.96 | -0.14 | 0.04 | 0.004 | W |  | 0.14 | 0.04 |  | degradation | AC POWER @ UPF | 0.6 | within limit | within limit
AC POWER @ UPF | 1500 | W | 1500.1 | 1500 | W | 1500.62 | -0.1 | -0.62 | -0.062 | W |  | 0.1 | 0.62 |  | degradation | AC POWER @ UPF | 1.2 | within limit | within limit
AC POWER @ UPF | 3000 | W | 3000.5 | 3000 | W | 3000.78 | -0.5 | -0.78 | -0.078 | W |  | 0.5 | 0.78 |  | degradation | AC POWER @ UPF | 2.4 | within limit | within limit
AC POWER @ UPF |  |  |  | 30 | W | 29.993 |  | 0.007 | 0.0007 | W |  |  | 0.007 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ UPF |  |  |  | 300 | W | 300.127 |  | -0.127 | -0.0127 | W |  |  | 0.127 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ UPF |  |  |  | 1500 | W | 1499.85 |  | 0.15 | 0.015 | W |  |  | 0.15 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ UPF |  |  |  | 3000 | W | 3000.8 |  | -0.8 | -0.08 | W |  |  | 0.8 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ UPF | 6000 | W | 5999.8 | 6000 | W | 6001.46 | 0.2 | -1.46 | -0.146 | W |  | 0.2 | 1.46 |  | improved |  | 4.8 | within limit | within limit
AC POWER @ UPF |  |  |  | 60 | W | 59.987 |  | 0.013 | 0.0013 | W |  |  | 0.013 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ UPF | 12000 | W | 12000.3 | 12000 | W | 12000.9 | -0.3 | -0.9 | -0.09 | W |  | 0.3 | 0.9 |  | degradation | AC POWER @ UPF | 9.6 | within limit | within limit
AC POWER @ 0.5PF(Lead) |  |  |  | 2.5 | W | 2.4996 |  | 0.0004 | 4e-05 | W |  |  | 0.0004 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lead) |  |  |  | 100 | W | 100.041 |  | -0.041 | -0.0041 | W |  |  | 0.041 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lead) |  |  |  | 7.5 | W | 7.5027 |  | -0.0027 | -0.00027 | W |  |  | 0.0027 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lead) | 75 | W | 75.03 | 75 | W | 75.008 | -0.03 | -0.008 | -0.0008 | W |  | 0.03 | 0.008 |  | degradation | AC POWER @ 0.5PF(Lead) | 0.06 | within limit | within limit
AC POWER @ 0.5PF(Lead) | 375 | W | 375.27 | 375 | W | 375.119 | -0.27 | -0.119 | -0.0119 | W |  | 0.27 | 0.119 |  | degradation | AC POWER @ 0.5PF(Lead) | 0.3 | within limit | within limit
AC POWER @ 0.5PF(Lead) | 750 | W | 750.4 | 750 | W | 750.028 | -0.4 | -0.028 | -0.0028 | W |  | 0.4 | 0.028 |  | degradation | AC POWER @ 0.5PF(Lead) | 0.6 | within limit | within limit
AC POWER @ 0.5PF(Lead) | 1500 | W | 1501.1 | 1500 | W | 1500.49 | -1.1 | -0.49 | -0.049 | W |  | 1.1 | 0.49 |  | degradation | AC POWER @ 0.5PF(Lead) | 1.2 | within limit | within limit
AC POWER @ 0.5PF(Lead) |  |  |  | 15 | W | 15.019 |  | -0.019 | -0.0019 | W |  |  | 0.019 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lead) |  |  |  | 150 | W | 150.035 |  | -0.035 | -0.0035 | W |  |  | 0.035 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lead) |  |  |  | 750 | W | 750.023 |  | -0.023 | -0.0023 | W |  |  | 0.023 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lead) | 1500 | W | 1501.36 | 1500 | W | 1500.35 | -1.36 | -0.35 | -0.035 | W |  | 1.36 | 0.35 |  | degradation | AC POWER @ 0.5PF(Lead) | 1.2 | out of limit | within limit
AC POWER @ 0.5PF(Lead) | 3000 | W | 3001.1 | 3000 | W | 3000.49 | -1.1 | -0.49 | -0.049 | W |  | 1.1 | 0.49 |  | degradation | AC POWER @ 0.5PF(Lead) | 2.4 | within limit | within limit
AC POWER @ 0.5PF(Lead) |  |  |  | 30 | W | 30.019 |  | -0.019 | -0.0019 | W |  |  | 0.019 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lead) | 6000 | W | 6002.5 | 6000 | W | 6001.51 | -2.5 | -1.51 | -0.151 | W |  | 2.5 | 1.51 |  | degradation | AC POWER @ 0.5PF(Lead) | 4.8 | within limit | within limit
AC POWER @ 0.5PF(Lag) |  |  |  | 2.5 | W | 2.4998 |  | 0.0002 | 2e-05 | W |  |  | 0.0002 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lag) |  |  |  | 100 | W | 100.059 |  | -0.059 | -0.0059 | W |  |  | 0.059 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lag) |  |  |  | 7.5 | W | 7.4989 |  | 0.0011 | 0.00011 | W |  |  | 0.0011 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lag) | 75 | W | 75.01 | 75 | W | 75.011 | -0.01 | -0.011 | -0.0011 | W |  | 0.01 | 0.011 |  | degradation | AC POWER @ 0.5PF(Lag) | 0.06 | within limit | within limit
AC POWER @ 0.5PF(Lag) | 375 | W | 375.07 | 375 | W | 375.067 | -0.07 | -0.067 | -0.0067 | W |  | 0.07 | 0.067 |  | degradation | AC POWER @ 0.5PF(Lag) | 0.3 | within limit | within limit
AC POWER @ 0.5PF(Lag) | 750 | W | 749.7 | 750 | W | 750.012 | 0.3 | -0.012 | -0.0012 | W |  | 0.3 | 0.012 |  | improved |  | 0.6 | within limit | within limit
AC POWER @ 0.5PF(Lag) | 1500 | W | 1499.9 | 1500 | W | 1500.2 | 0.1 | -0.2 | -0.02 | W |  | 0.1 | 0.2 |  | improved |  | 1.2 | within limit | within limit
AC POWER @ 0.5PF(Lag) |  |  |  | 15 | W | 15.006 |  | -0.006 | -0.0006 | W |  |  | 0.006 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lag) |  |  |  | 150 | W | 150.036 |  | -0.036 | -0.0036 | W |  |  | 0.036 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lag) |  |  |  | 750 | W | 750.048 |  | -0.048 | -0.0048 | W |  |  | 0.048 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lag) | 1500 | W | 1499.68 | 1500 | W | 1500.24 | 0.32 | -0.24 | -0.024 | W |  | 0.32 | 0.24 |  | improved |  | 1.2 | within limit | within limit
AC POWER @ 0.5PF(Lag) | 3000 | W | 3000.2 | 3000 | W | 3000.3 | -0.2 | -0.3 | -0.03 | W |  | 0.2 | 0.3 |  | degradation | AC POWER @ 0.5PF(Lag) | 2.4 | within limit | within limit
AC POWER @ 0.5PF(Lag) |  |  |  | 30 | W | 30.011 |  | -0.011 | -0.0011 | W |  |  | 0.011 |  | improved |  | 0 | out of limit | out of limit
AC POWER @ 0.5PF(Lag) | 6000 | W | 6000.1 | 6000 | W | 6001.25 | -0.1 | -1.25 | -0.125 | W |  | 0.1 | 1.25 |  | improved |  | 4.8 | within limit | within limit
Amplitude Vertical (DCV) | 1 | mV | 1.0294 | 1 | mV | 1.0006 | -0.0294 | -0.00059 | -5.9e-07 | V |  | 0.0294 | 0.00059 |  | degradation | Amplitude Vertical (DCV) | 2.5e-06 | out of limit | out of limit
Amplitude Vertical (DCV) | 10 | mV | 10.0349 | 10 | mV | 10.0005 | -0.0349 | -0.00048 | -4.8e-07 | V |  | 0.0349 | 0.00048 |  | degradation | Amplitude Vertical (DCV) | 2.5e-05 | out of limit | out of limit
Amplitude Vertical (DCV) |  |  |  | 100 | mV | 100.0002 |  | -0.0002 | -2e-07 | V |  |  | 0.0002 |  | improved |  | 0 | out of limit | out of limit
Amplitude Vertical (DCV) |  |  |  | 1 | V | 1 |  | -4.56e-05 | -4.56e-05 | V |  |  | 4.56e-05 |  | improved |  | 0 | out of limit | out of limit
Amplitude Vertical (DCV) |  |  |  | 10 | V | 10.0015 |  | -0.001516 | -0.001516 | V |  |  | 0.001516 |  | improved |  | 0 | out of limit | out of limit
Amplitude Vertical (DCV) |  |  |  | 30 | V | 29.9995 |  | 0.00051 | 0.00051 | V |  |  | 0.00051 |  | improved |  | 0 | out of limit | out of limit
Amplitude Vertical (ACV) |  |  |  | 2 | mV | 1.9985 |  | 0.0015 | 1.5e-06 | V |  |  | 0.0015 |  | improved |  | 0 | out of limit | out of limit
Amplitude Vertical (ACV) |  |  |  | 10 | mV | 9.9978 |  | 0.0022 | 2.2e-06 | V |  |  | 0.0022 |  | improved |  | 0 | out of limit | out of limit
Amplitude Vertical (ACV) |  |  |  | 100 | mV | 99.9753 |  | 0.0247 | 2.47e-05 | V |  |  | 0.0247 |  | improved |  | 0 | out of limit | out of limit
Amplitude Vertical (ACV) |  |  |  | 1 | V | 1 |  | -3.9e-05 | -3.9e-05 | V |  |  | 3.9e-05 |  | improved |  |  | within limit | out of limit
Amplitude Vertical (ACV) |  |  |  | 10 | V | 10.0004 |  | -0.00038 | -0.00038 | V |  |  | 0.00038 |  | improved |  |  | within limit | out of limit
Amplitude Vertical (ACV) |  |  |  | 100 | V | 100.0051 |  | -0.0051 | -0.0051 | V |  |  | 0.0051 |  | improved |  |  | within limit | out of limit
TIME | 2 | ns | 2 | 2 | ns | 2 | 9e-07 | 0 | 0 | s |  | 9e-07 | 0 |  | improved |  |  | out of limit | within limit
TIME | 10 | ns | 10 | 10 | ns | 10 | 4.2e-06 | 0 | 0 | s |  | 4.2e-06 | 0 |  | improved |  |  | out of limit | within limit
TIME | 100 | ns | 100 | 100 | ns | 100 | 4.11e-05 | 0 | 0 | s |  | 4.11e-05 | 0 |  | improved |  |  | out of limit | within limit
TIME | 1 | µs | 1 | 1 | µs | 1 | 4e-07 | 0 | 0 | s |  | 4e-07 | 0 |  | improved |  |  | out of limit | within limit
TIME | 10 | µs | 10 | 10 | µs | 10 | 4e-06 | 0 | 0 | s |  | 4e-06 | 0 |  | improved |  |  | out of limit | within limit
TIME | 100 | µs | 100 | 100 | µs | 100 | 4.06e-05 | 0 | 0 | s |  | 4.06e-05 | 0 |  | improved |  |  | out of limit | within limit
TIME | 1 | ms | 1 | 1 | ms | 1 | 4e-07 | -1e-07 | -1e-10 | s |  | 4e-07 | 1e-07 |  | improved |  |  | out of limit | out of limit
TIME | 10 | ms | 10 | 10 | ms | 10 | 4.2e-06 | -1e-06 | -1e-09 | s |  | 4.2e-06 | 1e-06 |  | improved |  |  | out of limit | out of limit
TIME | 100 | ms | 100 | 100 | ms | 100 | -1.19e-05 | -2e-05 | -2e-08 | s |  | 1.19e-05 | 2e-05 |  | degradation | TIME |  | out of limit | out of limit
TIME | 1000 | ms | 1000.0019 | 1000 | ms | 1000 | -0.0019 | 2e-05 | 2e-08 | s |  | 0.0019 | 2e-05 |  | degradation | TIME |  | out of limit | out of limit
Level/Bandwith (Frequency) | 10 | kHz | 10 | 10 | kHz | 10 | -4e-06 | -1e-06 | -0.001 | Hz |  | 4e-06 | 1e-06 |  | improved |  | 0.00025 | within limit | within limit
Level/Bandwith (Frequency) |  |  |  | 100 | kHz | 100 |  | -1e-05 | -0.01 | Hz |  |  | 1e-05 |  | improved |  | 0 | out of limit | out of limit
Level/Bandwith (Frequency) |  |  |  | 1 | MHz | 1 |  | -1e-07 | -0.1 | Hz |  |  | 1e-07 |  | improved |  | 0 | out of limit | out of limit
Level/Bandwith (Frequency) |  |  |  | 10 | MHz | 10 |  | -1e-06 | -1 | Hz |  |  | 1e-06 |  | improved |  | 0 | out of limit | out of limit
Level/Bandwith (Frequency) |  |  |  | 100 | MHz | 100 |  | -1e-05 | -10 | Hz |  |  | 1e-05 |  | improved |  | 0 | out of limit | out of limit
Level/Bandwith (Frequency) |  |  |  | 600 | MHz | 600.0001 |  | -6e-05 | -60 | Hz |  |  | 6e-05 |  | improved |  | 0 | out of limit | out of limit

## Sheet4
objective
we have 2021 data
we have permissible limit
from here we want to predict when the error will be neck to neck wth the permissible data
approach for work
train your model on 2019 data
time interval is known to you which is of 2 years
if your model's accuracy is good, apply the same model to predic the time interval where you have 2021 actual data and the permissible error limit
The result you obtain may put a question mark on its accuracy. You can then apply this model on 2019 data and 2021 data and show that the time interval is 2 years.
error 2019 | error 2021
error t-1 | error t
predicted time int = f (error t-1, error t)
predicted time interval = 2.01 years
actual interval = 2 years
predict time int = f(error 2021, permissible error)

## Sheet3
mV | 1000 | 3 | V
V | 1 | 0 | V
µA | 1000000 | 6 | A
mA | 1000 | 3 | A
A | 1 | 0 | A
Ω | 1 | 0 | Ω
kΩ | 0.001 | -3 | Ω
MΩ | 1e-06 | -6 | Ω
Hz | 1 | 0 | Hz
kHz | 0.001 | -3 | Hz
MHz | 1e-06 | -6 | Hz
°C | 1 | 0 | °C
nF | 1000000000 | 9 | F
PF | 1000000000000 | 12 | F
W | 10 | 1 | W
ns | 1000000000 | 9 | s
µs | 1000000 | 6 | s
ms | 1000 | 3 | s

## Sheet2
error = f(temperature point, type of thermocouple)